## Slide 1
FINAL PROJECT REVIEW
Measuring When a Brain Tumour
Segmentation Model Should Not Be Trusted
A multi-institution generalization, data-efficiency and uncertainty study on lower-grade glioma MRI
Abdul Azeez
Saripalli Sri Nandan
Dukka Chanakya Tej
AM.SC.U4AIE24049
AM.SC.U4AIE24043
AM.SC.U4AIE24070
B.Tech CSE (Artificial Intelligence)
Amrita Vishwa Vidyapeetham, Amritapuri
October 2026
Notes: Open by stating the reframe: we are not chasing a Dice score. We are asking when the model can be trusted.

## Slide 2
PROBLEM STATEMENT
A Dice score describes a population, not the scan in front of you
THE CONVENTIONAL RESULT
U-Net models reach human-comparable Dice on brain tumour MRI. This is a solved-looking problem, and a very common student project.
WHAT IT DOES NOT TELL YOU
A model correct on 82% of cases and confidently wrong on the rest is indistinguishable, at the point of use, from one that is always right.
Three questions, three controlled experiments, 24 trained models
RQ1
Does an unfamiliar hospital degrade accuracy - and if so, is the hospital really the cause?
RQ2
How much training data does this task need, and how small an effect can this dataset resolve?
RQ3
Can the model flag its own failures without ground truth, and where does that break down?
TrustSeg - Brain Tumour Segmentation
1
Notes: Emphasise: the contribution is the evaluation, not the architecture.

## Slide 3
DATASET AND SETUP
One public cohort that is secretly a five-institution dataset
LGG SEGMENTATION COHORT
110 patients · 3,929 FLAIR slices · 256×256
TCGA lower-grade glioma collection via TCIA, released by Buda et al. (2019)
Expert-verified binary FLAIR-abnormality masks
Every patient folder is named TCGA_<SITE>_<id>_<date> — the institution is in the filename, and is unused in prior work on this cohort
Site | DU | HT | CS | FG | EZ | Total
Patients | 45 | 34 | 16 | 14 | 1 | 110
EZ contributes one patient and stays in training throughout, so four sites are testable. The 45-vs-14 imbalance between DU and FG is what creates the confound in Experiment 1.
IDENTICAL CONFIGURATION ACROSS ALL 24 MODELS - ONLY THE SPLIT FILE CHANGES
Architecture
U-Net + ResNet-34 encoder,
ImageNet-pretrained, 24.4M params
Loss & optimiser
Dice + BCE · Adam 1e-4 ·
ReduceLROnPlateau · early stopping
Augmentation
Geometric only — intensity augmentation
would erase the effect under study
Hardware
RTX 4050, 6 GB VRAM · batch 4 ·
mixed precision (FP16)
TrustSeg - Brain Tumour Segmentation
2
Notes: The geometric-only augmentation choice is deliberate and defensible. Expect a question.

## Slide 4
EXPERIMENT 1
Leave-one-hospital-out has a confound, so we controlled for it
The confound: excluding DU removes 45 of 110 patients. Excluding FG removes 14. Site identity and training-set size change together, so any degradation is attributable to either.
A
Baseline
Patient-level random split across all institutions — the conventional protocol.
1 model
B
Leave-one-hospital-out
One institution excluded from train and validation entirely, used only as the test set.
4 models
C
Matched control
Same train/val/test patient counts as each LOHO split, but test patients drawn from all institutions.
4 models
Site effect = Dice(B) − Dice(C) Volume effect = Dice(C) − Dice(A)
TrustSeg - Brain Tumour Segmentation
3
Notes: Arm C is the methodological contribution — it is absent from prior cross-institutional work.

## Slide 5
EXPERIMENT 1 — RESULT
The hospital is not the cause
Site | Test pts | Dice LOHO | Dice control | Site effect | Volume effect | p
DU | 45 | 0.7052 | 0.6915 | +0.0137 | −0.1262 | 0.784
HT | 34 | 0.7420 | 0.7268 | +0.0152 | −0.0908 | 0.213
CS | 16 | 0.7173 | 0.8228 | −0.1055 | +0.0052 | 0.337
FG | 14 | 0.7979 | 0.7291 | +0.0688 | −0.0885 | 0.260
Mean | — | 0.7406 | 0.7426 | −0.0020 | −0.0751 | —
SITE EFFECT
−0.0020
Indistinguishable from zero. Three of four sites scored HIGHER unseen.
VOLUME EFFECT
−0.0751
Accounts for essentially the entire apparent 0.077 gap.
SIGNIFICANCE
p ≥ 0.21
No site reaches significance (Mann-Whitney U, per-patient Dice).
The clearest single refutation: Control — DU excluded no institution at all, yet scored 0.6915 — lower than LOHO — DU at 0.7052.
Baseline 0.8176 reproduces the 0.82 reported by Buda et al. (2019) on this cohort, confirming the backbone is correctly trained.
TrustSeg - Brain Tumour Segmentation
4
Notes: Point at Control-DU vs LOHO-DU. That one comparison does the work.

## Slide 6
EXPERIMENT 2
How much data does this task actually need?
Chart: Test Dice vs number of training patients
| 15 | 30 | 45 | 60 | 75
Measured (3 seeds, fixed test set) | 0.6681 | 0.7262 | 0.7496 | 0.7862 | 0.7976
FITTED CURVE
Dice = 0.918 − 13.6/(n + 39.8)
Projected ceiling 0.918. Diminishing returns beyond ~96 training patients — more than this cohort contains, so the model is still data-limited.
NOISE FLOOR
± 0.057 Dice
Same size, same test set, different patients drawn. Effects below this cannot be resolved.
15 models · 5 sizes × 3 seeds
The 3 seeds matter: they separate WHICH patients were drawn from HOW MANY. That is what makes the noise floor a measurement rather than a guess.
TrustSeg - Brain Tumour Segmentation
5
Notes: The curve is the positive, quantitative result that the negative finding in Experiment 1 pointed toward.

## Slide 7
EXPERIMENT 2 — THE DECISIVE COMPARISON
Neither explanation survives
0.1313
Observed spread across the 8 cross-site models
0.0319
What the learning curve predicts over that training-size range
0.0566
Training-side noise floor (2 SD)
All eight cross-site models trained on 55–82 patients — a region where the curve moves by only 0.032. Their scores span 0.131. Neither training size nor seed noise can explain a spread four times larger than either.
THE FINDING
Per-institution validation on fewer than ~50 patients cannot detect domain shift of the magnitude the literature reports. Test-set sampling variance exceeds the effect by roughly 4×. A study this size that reports a site effect is most likely reporting noise.
TrustSeg - Brain Tumour Segmentation
6
Notes: This is the strongest methodological claim in the project. Deliver it slowly.

## Slide 8
EXPERIMENT 3
Can the model flag its own failures?
WHY NOT MC DROPOUT
MC Dropout needs dropout layers active at inference. The segmentation-models-pytorch U-Net has none, so it cannot run on these checkpoints without retraining the architecture.
TEST-TIME AUGMENTATION INSTEAD
Each slice is predicted 8 times — the four 90° rotations, each with and without a mirror — and every prediction is mapped back. All 8 should coincide. Their disagreement is the uncertainty.
Why this is the honest choice here: the model was TRAINED with flips and 90° rotations, so it has every reason to be invariant to them. Disagreement is genuine instability, not an unfamiliar input. We unit-tested that all 8 transforms invert exactly — without that the whole analysis would be invalid.
Uncertainty score | rho vs Dice (tumour slices) | rho vs Dice (all slices) | Used for
Foreground entropy | −0.862 | +0.606 | offline analysis
View disagreement | −0.628 | −0.885 | the live system
Mean entropy | −0.108 | −0.778 | —
Boundary entropy | −0.116 | −0.883 | —
Foreground entropy ranks better where a tumour is present, but divides by predicted area so its sign inverts on empty slices. View disagreement is weaker but holds its sign everywhere — which is what deployment needs.
TrustSeg - Brain Tumour Segmentation
7
Notes: Two scores, two jobs. Be ready to explain why we did not just pick the bigger number.

## Slide 9
EXPERIMENT 3 — RESULT
Referring the least-confident 20% raises Dice from 0.81 to 0.90
Chart: Dice on retained slices as uncertain cases are referred
| 0 | 10 | 20 | 30 | 40 | 50
Foreground entropy | 0.8075 | 0.8817 | 0.9014 | 0.9176 | 0.9262 | 0.9344
View disagreement | 0.8075 | 0.8431 | 0.8603 | 0.8673 | 0.8679 | 0.8644
Deployment bands, validated on 715 slices from 25 patients
Verdict band | Slices | Dice
CONFIDENT | 128 | 0.938
REVIEW | 46 | 0.847
HIGH RISK | 28 | 0.656
MISSED FINDING | 11 | 0.000
WHY THIS MATTERS
Dice falls and uncertainty rises monotonically across the bands. That is what makes the traffic light meaningful rather than cosmetic — it is measured on slices chosen independently of the test split.
Spearman rho = −0.86 between uncertainty and realised Dice, on 279 tumour-bearing test slices.
TrustSeg - Brain Tumour Segmentation
8
Notes: This is the positive result and the demo feature. Lead with the 0.81 to 0.90 number.

## Slide 10
EXPERIMENT 3 — THE LIMIT
Where the uncertainty mechanism fails, and why it must
Of 224 tumour-bearing slices, 22 (9.8%) are predicted completely empty. They split cleanly in two:
11 SLICES — CAUGHT
Individual views DO find the abnormality, but the averaged probability never crosses 0.5, so the mask comes out empty while the 8 views are split. The most uncertain state the model can occupy. The system reports these as a possible missed finding.
11 SLICES — NOT CAUGHT
Disagreement is exactly zero. All 8 views unanimously find nothing, and all 8 are wrong. There is no signal to detect. This is roughly 5% of tumour-bearing slices.
Disagreement-based uncertainty cannot detect a failure the model is uniformly confident about. This is the structural limit of the method, not an implementation defect — and we measured exactly where it lies.
The cost of checking uncertainty first: 5.5% of tumour-free slices are flagged high risk. A deliberate trade — a false alarm costs a second look, a missed tumour costs more.
TrustSeg - Brain Tumour Segmentation
9
Notes: Raise this yourself. If an examiner finds it first it reads as a hole; raised first it reads as rigour.

## Slide 11
THE SYSTEM
TrustSeg — segmentation that reports its own confidence
Input
A FLAIR slice is uploaded and resized to 256×256.
8-view pass
The slice is predicted 8 times under the D4 symmetry group and every prediction mapped back.
Output
Tumour mask, per-pixel uncertainty heat map, area, hemisphere, how many of the 8 views found an abnormality.
Verdict
A calibrated band: CONFIDENT / REVIEW / HIGH RISK / POSSIBLE MISSED FINDING.
Three decisions that privilege honesty over appearance
—
Uncertainty is checked BEFORE predicted area. An empty mask with a split vote is reported as a possible missed finding, never as a clean negative.
—
Area is in pixels unless spacing is supplied. The source TIFs carry no pixel spacing, so mm² would be fabricated.
—
Hemisphere comes from the image midline. Labelled as such. Lobe-level localisation would need atlas registration, which we do not perform.
TrustSeg - Brain Tumour Segmentation
10
Notes: Switch to the live demo here. Four slices: confident, review, high risk, missed finding.

## Slide 12
LIMITATIONS
Stated first, so they are evidence of rigour rather than holes
Statistical power
With 14–45 test patients per site we can exclude a large site effect but not a small one. No effect is DETECTABLE at this sample size — that is not the same as none existing.
Attribution
Site, scanner vendor and protocol are inseparable in TCGA metadata. Even a confirmed effect could not be traced to a specific scanner property.
The blind spot
~5% of tumour-bearing slices are missed with zero disagreement and cannot be flagged. Fixing this needs an orthogonal uncertainty source, such as a model ensemble.
False alarms
5.5% of tumour-free slices are flagged high risk, clustered by patient rather than spread evenly.
Scope
FLAIR only, 2D slice-wise, lower-grade glioma. Tumour-type classification is impossible here — every patient shares one diagnosis.
Every one of these is a measurement, not an estimate.
TrustSeg - Brain Tumour Segmentation
11
Notes: Volunteering limitations pre-empts the obvious questions and demonstrates understanding.

## Slide 13
CONCLUSION
What 24 trained models established
1
The hospital is not the cause
Site effect −0.0020, no significance. Reproduced independently by the learning-curve residuals (−0.0023).
2
Per-site validation under ~50 patients is under-powered
Sampling variance 0.131 against an explainable effect of 0.032 — a factor of four.
3
The data requirement is quantified
Ceiling 0.918 Dice, diminishing returns beyond ~96 patients — more than this cohort holds.
4
The model can flag its own failures
rho = −0.86; referring the least-confident 20% raises Dice 0.81 → 0.90. Its 5% blind spot is measured too.
The contribution is not a higher Dice score and not a new architecture. It is a reproducible protocol for deciding whether a reported generalization result is real, and which individual predictions to trust.
TrustSeg - Brain Tumour Segmentation
12
Notes: Close on the framing sentence. That is the answer to "what did you actually contribute".

## Slide 14
REFERENCES
[1] Ronneberger, Fischer & Brox. "U-Net: Convolutional Networks for Biomedical Image Segmentation." MICCAI, 2015.
[2] He, Zhang, Ren & Sun. "Deep Residual Learning for Image Recognition." IEEE CVPR, 2016.
[3] Isensee et al. "nnU-Net: a self-configuring method for deep learning-based biomedical image segmentation." Nature Methods, 18, 2021.
[4] Zech et al. "Variable generalization performance of a deep learning model to detect pneumonia in chest radiographs." PLOS Medicine, 15(11), 2018.
[5] AlBadawy, Saha & Mazurowski. "Deep learning for segmentation of brain tumors: Impact of cross-institutional training and testing." Medical Physics, 45(3), 2018.
[6] Karani, Chaitanya, Baumgartner & Konukoglu. "A Lifelong Learning Approach to Brain MR Segmentation Across Scanners and Protocols." MICCAI, 2018.
[7] Mårtensson et al. "The reliability of a deep learning model in clinical out-of-distribution MRI data." Medical Image Analysis, 66, 2020.
[8] Gal & Ghahramani. "Dropout as a Bayesian Approximation." ICML, 2016.
[9] Kendall & Gal. "What Uncertainties Do We Need in Bayesian Deep Learning for Computer Vision?" NeurIPS, 2017.
[10] Mehta et al. "QU-BraTS: MICCAI BraTS 2020 Challenge on Quantifying Uncertainty in Brain Tumor Segmentation." MELBA, 2022.
[11] Buda, Saha & Mazurowski. "Association of genomic subtypes of lower-grade gliomas with shape features automatically extracted by a deep learning algorithm." Computers in Biology and Medicine, 109, 2019. [dataset]
[12] Menze et al. "The Multimodal Brain Tumor Image Segmentation Benchmark (BRATS)." IEEE TMI, 34(10), 2015.
Dataset: LGG Segmentation Dataset (TCGA-LGG, The Cancer Imaging Archive), released with reference [11].
TrustSeg - Brain Tumour Segmentation
13

## Slide 15
Thank you
Questions
Abdul Azeez · Saripalli Sri Nandan · Dukka Chanakya Tej
B.Tech CSE (AI), Amrita Vishwa Vidyapeetham, Amritapuri